## Slide 1: Preparing Calibration Service Procedures and Validation Methods in Support of CMCs
Marina Gertsvolf
Jan 28, 2014
2015 SIM TFWG Workshop

## Slide 2: CMC review example
Recent CMS from COOMET_TF_KZ

## Slide 3: NRC CMC in Time and Frequency
General Frequency Source Calibration – Procedure Missing (external audit finding)

## Slide 4: Calibration Approach
Method
  Frequency source  frequency counter or time interval counter
    Procedure FTS_FREQMEAS
  frequency counter: calibration procedure missing
    Procedure FTS-FREQ_CNTR_CAL
  time interval counter : verified for the use at 1PPS only via CCTF-K001.UTC

## Slide 5: Calibration Procedure: FTS-FREQMEAS

## Slide 6: Calibration Procedure: FTS-FREQ_CNTR_CAL

## Slide 7: Calibration Validation
Procedure:
    Frequency source  frequency counter or time interval counter
Validation:
  Known frequency source  Calibrated frequency counter or time interval counter
  Over all frequency ranges and with the worst case slope
    100 ns rise time
    1 Vpp
    0.1 Hz
    1 hr averaging

## Slide 8: Calibration Validation: Setup
Time Interval Counter
HP 5370B
Measurement Automation Switch Box
NRC R.P.
UTC(NRC)
1PPS
TTL
Start with the system used in the Key Comparison CCTF-K001.UTC
H maser or Cs clock
5MHz
sine
time base

## Slide 9: Results
HP-5370B measurement results example

## Slide 10: Results
UTC - UTC(NRC) from Circular T

## Slide 11: Calibration Validation: Setup
Generate custom signals with CMC specifications
H maser or Cs clock
5MHz
sine
Frequency Synthesizer
Anritsu MG 3641A
Stanford Research DS245
Frequency Divider
NRC CAL DIV W.P.
5 - 100MHz
sine
nPPS
TTL
2.32Vpp @50Ω
time base

## Slide 12: Calibration Validation: Setup
Use K001.UTC system for custom signals calibration
Time Interval Counter
HP 5370B
Measurement Automation Switch Box
NRC R.P.
UTC(NRC)
1PPS
TTL
H maser or Cs clock
5MHz
sine
5MHz
sine
Frequency Synthesizer
Anritsu MG 3641A
Stanford Research DS245
Frequency Divider
NRC CAL DIV W.P.
5 - 100MHz
sine
nPPS
TTL

## Slide 13: Frequency Measurements with HP system
MG 3641A output 50MHz CAL DIV output 1E-7 HP 5370B measurement | Allan Deviation
CAL DIV output 1E-8 Pendulum CNT-91 measurement | MG 3641A output 50MHz Pendulum CNT-91 measurement
MG 3641A output 50MHz CAL DIV output 1E-7 HP 5370B measurement | Allan Deviation

## Slide 14: Frequency Measurements with HP system part 2
MG 3641A output 3991680.10 Hz CAL DIV output 1E-7 HP 5370B measurement | MG 3641A output 3991680.10 Hz CNT-91 measurement
 | Pendulum CNT-91 measurement MG 3641A output 50MHz ; CAL DIV output 1E-6 ;
MG 3641A output 3991680.10 Hz CAL DIV output 1E-7 HP 5370B measurement | MG 3641A output 3991680.10 Hz CNT-91 measurement

## Slide 15: Time Interval counter readings at non-integer frequencies
where
= fractional frequency offset from
= initial offset
= number of seconds = START pulse count
Simulations show that this function can fit measurement results with noise levels of up to in time jitter.
STOP
START
= 1s = TUTC(NRC)
1/f0 = T0
Δt0
Δtn

## Slide 16: Frequency Measurements with HP system part 3
MG 3641A output 3991680.10 Hz CAL DIV output 1E-7 HP 5370B measurements UNWRAPPED | Allan Deviation
PROBLEM: The HP 5370B data acquisition software generates one output value calculated from an average of 11 points around the nominal time stamp. The accuracy of the timestamp value becomes critical when the DUT frequency is not a multiple of Hz and over 11 measurements the time interval varies over ns or more. As a result, when the DUT frequency, after being divided and delivered to the HP 5370B is not an integer number of Hz (± 1.0e-9), the averaging over 11 measurements leads to very high modulation in the reported readings.

## Slide 17: Calibration Validation: Setup
Use Calibrated Frequency / Time Interval Counter system for custom signals calibration
H maser or Cs clock
5MHz
sine
5MHz
sine
Frequency Synthesizer
Anritsu MG 3641A
Stanford Research DS245
Frequency Divider
NRC CAL DIV W.P.
5 - 100MHz
sine
nPPS
TTL
UTC(NRC)
1PPS
TTL
Frequency / Time Interval Counter
Pendulum CNT-91

## Slide 18: Calibration Validation
New software for data acquisition with CNT-91

## Slide 19: Frequency Measurements with CNT system, CNT measurements validation against HP system (traceable to UTC via CCTF-K001.UTC
MG 3641A output 100MHz CAL DIV output 1E-7  HP 5370B (blue) CAL DIV output 1E-8  CNT-91(red) | Allan Deviation

## Slide 20: Frequency Measurements with CNT system, Measurements of non-integer frequency values
MG 3641A output 47.9001599 MHz CAL DIV output 1E-7 CNT-91 measurement | Allan Deviation

## Slide 21: Calibration Validation
Low Pass filter for simulating 100ns rise time
  Connected between the Frequency Divider and the Counter
Attenuator for simulating 1Vpp
  Connected after the filter
V in
V out
R = 3 Ω
c = 2.2 pF
TTL
rise time 120 ns
measured with 500MHz BW TDS 3054B oscilloscope

## Slide 22: Frequency Measurements with 1Vpp, 120ns rise time
MG 3641A output 100MHz CAL DIV output 1E-8 CNT-91 measurement | Allan Deviation

## Slide 23: Calibration Validation: Frequency Counter Calibration using Known Frequencies
Validate Frequency Counter readings using known frequencies
H maser or Cs clock
Frequency multiplier
Spectra Dynamics FS100-RM
5MHz
sine
10,40,90 MHz
sine
2.32Vpp @50Ω
5MHz
sine
Frequency / Time Interval Counter
Pendulum CNT-91

## Slide 24: Frequency Counter Calibration using Known Frequencies
The CNT-91 was used in a frequency counter mode
FS100-RM Frequency Output 2.32Vpp @ 50Ω [MHz] | CNT-91 Reading A: AC, _¦¯, 50Ω, 1x, Man, 0V MeasTime 1s, # samples 100 [MHz]; | CNT-91 Reading A: AC, _¦¯, 50Ω, 1x, Man, 0V MeasTime 10s, # samples 100 [MHz];
10 | 10; σy = 3.5e-11 | 10; σy = 2.5e-12
40 | 40; σy = 0.75e-11 | 40; σy = 1.5e-12
90 | 90; σy = 0.75e-11 | 90; σy = 1.5e-12

## Slide 25: Summary
The validation had verified the measurement capability and the traceability to the levels exceeding CMC specifications
The frequency counter calibration procedure was developed and the counter was calibrated

## Slide 26
26
Thank you
Dr. Marina Gertsvolf
Tel: 613-998-1311
marina.gertsvolf@nrc-cnrc.gc.ca
www.nrc-cnrc.gc.ca